## Sheet1
S. No. | ITI Specification | TENDER REQUIREMENT | Clarification
|  |  | Thermo | Waters | Sciex | Agilent | Shimadzu | Recommendation by members of BEC
1.1.1 | Electrospray Ionization Source – ESI | a. Desolvation temperature – 500 ºC or better | No comments | No comments | No comments | Desolvation temperature 400 ºC for ESI and 500 for APCI or better. (Software screenshot data can be provided as proof) | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
1.2.4 | Sensitivity # | Note: The below sensitivity and IDL must be mentioned in the product specification sheet published from the date of launch of the quoted model or should be supported with chromatographic evidence generated in one of the principal’s demo laboratories. | This clause will open opportunity for vendors to submit their demo lab data generated locally which are not standardized and will not be reproducable at customer site. | No comments | No comments | No comments | No comments | No changes required in the published specification
|  | No revised S/N specification will be accepted. Any deviation will lead to disqualification. |  | No comments | No comments | No comments | No comments | No changes required in the published specification
1.2.4.1 | ESI Sensitivity | a. Sensitivity of the instruments: In ESI +ve : 1 pg/μL of Reserpine, sensitivity should be 1,000,000:1 or better | Having better IDL has more analytical benefits for the system.1. Superior Chromatographic separation2. Higher sensitivity of the instrument3. Better S/N ratio4. Reduce background contamination, Hence we recommend to keep. | a. Sensitivity of the instruments: In ESI +ve :1pg/ul of Reserpine Sensitivity Should be 5,000,000:1 or better | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
|  | b. Sensitivity of the instruments: In ESI -ve : 1 pg/μL of Chloramphenicol sensitivity should be 1,000,000:1 or better |  | b. Sensitivity of the instruments: In ESI +ve :1pg/ul of Chloramphenicol Sensitivity Should be 5,000,000:1 or better | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
|  | c.IDL ESI+ve for reserpine should be ≤ 0.25 fg | IDL ESI+ve for reserpine should be ≤ 0.25 fg | c. IDL ESI +ve for Reserpine should be < 0.3 fg | IDL ESI+ve for reserpine should be ≤ 0.28 fg 1 d. | Lab data for 0.25fg or IDL ESI+ve for reserpine should be <0.3 fg | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
|  | d.IDL ESI-ve for chloramphenicol should be ≤0.25 fg | IDL ESI-ve for chloramphenicol should be ≤0.25 fg | No comment | IDL ESI-ve for chloramphenicol should be ≤0.25 fg d. | IDL ESI-ve for chloramphenicol should be <0.3 fg | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
1.2.4.2 | APCI Sensitivity | a.Sensitivity of the instruments: In APCI +ve : 200 fg/μL of Reserpine sensitivity should be 100,000:1 or better | No comments | request to be removed | Sensitivity of the instruments: In APCI +ve : 10 pg/μL of Reserpine sensitivity should be 100,000:1 or better | Sensitivity of the instruments: In APCI +ve: 200 fg/µL of Reserpine sensitivity should be 100,000:1 or better (Lab data generated is acceptable) Request to accept lab generated data | we would like to hereby confirm that supporting data and confirmations on S/N ratio for APCI (-) ve is not available in any technical literature in terms of Chloramphenicol. Therefore we would like to request to reconsider the mentioning of APCI negative sensitivity in terms of Chloramphenicol | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends no changes to the published specification
|  | b.Sensitivity of the instruments: In APCI -ve: 200 fg/μL of Chloramphenicol sensitivity should be 100,000:1 or better | As discussed please remove the APCI negative sensitivity. There are no universal reference compound for APCI negative sensitivity. | request to be removed | Please remove this Point |  | we would like to hereby confirm that supporting data and confirmations on S/N ratio for APCI (-) ve is not available in any technical literature in terms of Chloramphenicol. Therefore we would like to request to reconsider the mentioning of APCI negative sensitivity in terms of Chloramphenicol | Since unanimous concern was raised against the specification required by ITI, BEC recommends to change the published specification by removing the requirement for this clause on sensitivity of APCI in negative polarity
1.4 | a. Quadrupole Analyzer: | a. Mass Range - The quadrupole mass range 5 – 2000 m/z or better. | Better to keep the quadrupole mass range 5 – 2000 m/z or better considering its future application availability. | No comments | No comments | No comments | No comments | No changes required in the published specification
| b. Mass Resolution # | i.Reserpine - Resolution should be up to ≤0.25 Da FWMH for better selectivity across the entire mass range | We recommend to keep ≤0.25 Da FWMH resolution level as it provides | i. Reserpine- Resolution should be up to < 0.5 FWHM and automatically adjustable for better sensitivity across the mass range | Reserpine - Resolution should be up to ≤0.7 Da FWMH for better selectivity across the entire mass range | Resolution should be 0.4 to 0.7 Da FWMH range. | we would like to request to revise the specifications as “ The quoted LC MS/MS system must be capable of facilitating a mass resolution range between 0.4 to 0.7” | Since unanimous concern was not raised against the specification required by ITI, BEC recommends to not to change the published specification
1. Better mass selectivity
2. Reduce chemical noise
3. Better selectivity in MRM transitions
|  | ii. Chloramphenicol - Resolution should be ≤0.25 Da FWMH for better selectivity across the entire mass range | 4. Improve quantitation and accuracy | ii. Chloramphenicol - Resolution should be up to < 0.5 FWHM and automatically adjustable for better sensitivity across the mass range | Chloramphenicol - Resolution should be ≤0.7 Da FWMH for better selectivity across the entire mass range
| d. Polarity Switching Time | +ve / -ve polarity switching time within the source should be 5 ms or better | We recommend to keep the lower +ve / -ve polarity switching time as 5 ms or better to enable a single run for mixed analytes (positive and negative ions in one LC run) | Polarity Switching time - +ve/-ve polarity switching time should be 15 ms or better | No comment | Polarity Switching Time +ve / -ve polarity switching time within the source should be 25 ms or better | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
1. Enhance capability for multi residue fast LC method
2. Higher duty cycle efficiency
3. More data point per Chromatographic peak
1.4. | f. Selectivity | II. The system should have an open architecture/ open source spectral libraries which facilitate the mass spectral comparison in mycotoxins, veterinary antibiotic residues, and pesticide residue analysis. (Users should have the ability to add new records to the library) | No comments | No comments | II. The system should have a Latest NIIST 2023 libraries having more than 30000 compounds which facilitate the mass spectral comparison in mycotoxins, veterinary antibiotic residues, and pesticide residue analysis. (Users should have the ability to add new records to the library) | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
1.4 | e. Scan Speed - | Should have a scan speed of 12000 amu/sec or better | No comments | Scan Speed -Should have a scan speed of 20000 amu/sec or better | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
1.7 | Vacuum System | i. Should have minimum maintenance and utility | We request to increase the noise level to 60dBA | All agreed during the pre-bid meeting |  |  |  | Since unanimous concern was raised against the specification required by ITI, BEC recommends to clearly mention the specification required by ITI as follows; "ii. Noise level should be lower than the recommended for a standard laboratory of 60 dBA"
ii. Noise level should be lower than the recommended for a standard laboratory of 55 dBA as given in 1991 HVAC Applications
2 | Ultra- Performance Liquid Chromatography System/Ultra High-Performance Liquid Chromatography System | a.The UHPLC system and the MS should be able to be controlled individually by software. | It is not recommend to use UHPLC as a front end LC for stand along applications. Hence we suggest to change the specifications as "The UHPLC system and the MS should be able to be controlled by a single software | No comments | No comments | No comments | No comments | Since unanimous concern was not raised against the specification required by ITI, BEC recommends to not to change the published specification
2.1 | Pump | Operating flow range should be 0.001- 3.000 mL/min or better with 1 μL increments. | No comments | Operating Flow range should be 0.001 to 2 ml/min with 1 ul increments | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
|  | 8.Usable solvent types should include both organic and aqueous solutions and it should work at full scale of pH range of 1 to 14. | We request to change the pH range 2 to 12. pH 1 to 14 is not an universal range. Hence we request to change the pH range 2 to 12 | No comments | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
2.1 |  | 7.Gradient concentration accuracy ± 0.15% or better | No comments | No comments | No comments | No comments | We would like to confirm whether it is ok to have a gradient concentration accuracy of 0.5% if the quoted system and configuration can achieve LOD/LOQ levels specified for target analytes | Since unanimous concern was not raised against the specification required by ITI, BEC recommends to not to change the published specification
2.2 | Autosampler | 100 or better vial position in sample tray ( 2ml vial) | No comments | 96 or better vial position in sample tray ( 2ml vial) | No comments | No comments | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
2.2 | Autosampler, carryover | The carryover must be ≤ 0.002% with caffeine or better. | No comments | No comments | No comments | The carryover must be = 0.002% with caffeine or better (lab data on caffeine) Agilent measures carryover with more stringent criteria and stickier compound- Chlorhexidine – (0.004%) but if needed we can submit the data on caffeine which is 0.0007% which is far better. | No comments | Since unanimous concern was not raised against changing the specification required by ITI, BEC recommends to not to change the published specification
4 | System Controller and Operating System | 10. Provision to install the data processing software in an additional 5 nos of PCs free of charge | Please change the spec to "Provision to install the data processing software in an additional 4 nos of PCs free of charge" | No comments | Specify the Number of Licensed Processing Software needed. | No comments | No comments | Since unanimous concern was raised against the specification required by ITI, BEC recommends to clearly mention the specification required by ITI as follows;" Provision to install the data processing licenced software in an additional 4 nos of PCs free of charge (software only)."
8 | Installation and Commissioning | [5 empty] | Under 8 Installation/ Commissioning – 7 iii)- Eight Nitro furan and respective metabolites (LOD should be less than 0.5 µg/kg in matrix) in poultry meat samples- We would like to confirm the names of eight nitro furan and respective metabolites mentioned under this & under iv) Screening of at least 700 Pesticide residues in a single run with polarity switching (instrument detection limit should be 0.5 ppb in the matrix- Black Tea) ∆- We would like to know following; | iii) Furazolidone & 3-amino-2-oxazolidinone (AOZ), Furaltadone & 3-amino-5-morpholino-2-oxazolidinone (AMOZ), Nitrofurantoin & 1-aminohydantoin (AHD), and Ntrofurazone & semicarbozide (SEM)
[7 empty] | - What is the expected recovery ? | All are expected to be within the limits as required and recommended by the latest SANTE guidline on method validation requirements for pesticide residue analysis
-Are we expecting recovery for all pesticides compounds to be the same?
-What is the expected linearity are we looking at R2 >/ to 0.99 or 0.999?